## Slide 1
Global Business Dept.
Hanyoung Nux
Common Error & Troubleshooting
for Controller

## Slide 2: Contents
1. Temperature Humidity Controller
TD500
TH500
2. Programmable Temperature Controller
NP200
PX series
TH300
TH510
3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
4. Digital Temperature Controller
HY-48D
HY-72D
HY-8000
HY-8200S
5. Temperature Indicator
BK6
RT9
6. Counter/Timer
GE/GF series

## Slide 3: 1. Temperature Humidity Controller
TD500
TH300
TH500
TH510
1. Temperature Humidity Controller
Error Code |  | Error Description |  | Troubleshooting
- |  | Administrator Password(Default) |  | 0 or 1030
- |  | Sensor Error and DI Error Message |  | Check DI setting and respond to DI Error properly
D/I **** ERROR** |  | An abnormal signal is detected in the equipment from connected DI. |  | Please check the connected terminal to DI.

## Slide 4: 2. Programmable Temperature Controller
NP200
PX series
2. Programmable Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
B.OUT |  |  | - Thermocouple or RTD sensor disconnection or wiring problem - PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring with polarity * If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
SYS.ER |  |  | - The internal parameter error - When the input side is affected by noise - System error in EEProm |  | Please contact Hanyoung Nux.
ADC.ER |  |  | - Main board problem |  | After thorough inspection, repair possibility will be decided.

## Slide 5: 2. Programmable Temperature Controller
NP200
PX series
Error Code |  |  | Error Description |  | Troubleshooting
B.OUT |  |  | - Thermocouple or RTD sensor disconnection or wiring problem - PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring with polarity * If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
2. Programmable Temperature Controller

## Slide 6: 3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
B.OUT |  |  | - Thermocouple or RTD sensor disconnection or wiring problem - PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring with polarity * If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
E.ADC |  |  | Analog-to-digital ADC converter problem Input or power side may be affected by noise |  | 1. Reboot 2. If still error displayed, please check the wires if there is effective noise. 3. If the noise affect the controller, the controller should calibrate again. 4. But if there is noise closely, first remove the noise or prevent the noise.

## Slide 7: 3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OVR / -OVR |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring

## Slide 8: 3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OOOO |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Press UP & Down key for a while 2. Check SL5 = 0, SL7 = maximum value, SL8 = minimum value setting 3. Check if wiring is tightened 4. If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
UUUU |  |  | PV exceeding the maximum range of temperature |  | 1. Press UP & Down key for a while 2. Check SL5 = 0, SL7 = maximum value, SL8 = minimum value setting 3. Check if wiring is tightened 4. If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
SYS.E |  |  | - The internal parameter error - When the input side is affected by noise - System error in EEProm |  | Please contact Hanyoung Nux.
ADE,DDE |  |  | 250Ω resistors or internal AC/ DC converter problems in 4-20mA input |  | 1. Reboot 2. Remove the resistor 3. Remove the sensor 4. Check the error code and follow the corresponding troubleshooting

## Slide 9: 3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OVR / -OVR |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring

## Slide 10: 3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OVR / -OVR |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring
B.OUT |  |  | - Thermocouple or RTD sensor disconnection or wiring problem - PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring with polarity * If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
SYS.E |  |  | - The internal parameter error - When the input side is affected by noise - System error in EEProm |  | Please contact Hanyoung Nux.
ADE,DDE |  |  | 250Ω resistors or internal AC/ DC converter problems in 4-20mA input |  | 1. Reboot 2. Remove the resistor 3. Remove the sensor 4. Check the error code and follow the corresponding troubleshooting again

## Slide 11: 3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OOOO |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Press UP & Down key for a while 2. Check SL5 = 0, SL7 = maximum value, SL8 = minimum value setting 3. Check if wiring is tightened 4. If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
UUUU |  |  | PV exceeding the maximum range of temperature |  | 1. Press UP & Down key for a while 2. Check SL5 = 0, SL7 = maximum value, SL8 = minimum value setting 3. Check if wiring is tightened 4. If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
SYS.E |  |  | - The internal parameter error - When the input side is affected by noise - System error in EEProm |  | Please contact Hanyoung Nux.
ADE,DDE |  |  | 250Ω resistors or internal AC/ DC converter problems in 4-20mA input |  | 1. Reboot 2. Remove the resistor 3. Remove the sensor 4. Check the error code and follow the corresponding troubleshooting
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller

## Slide 12: 3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OVR / -OVR |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring
B.OUT |  |  | - Thermocouple or RTD sensor disconnection or wiring problem - PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring with polarity * If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
KSTP |  |  | Without DI option - Displayed when STOP |  | RUN↔STOP can be changed by pressing a while “ <</R/S” button on the operation screen
KSTP |  |  | When there is DI option - Display when stop operation by “<</R/S” button |  | 운전RUN↔STOP can be changed by pressing a while “ <</R/S” button on the operation screen

## Slide 13: 3. Digital PID Temperature Controller
AX series
BR6A
DX series
ED6
HX series
KX☐N series
MC9 Series
NX series
3. Digital PID Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OVR / -OVR |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring
B.OUT |  |  | - Thermocouple or RTD sensor disconnection or wiring problem - PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring with polarity * If thermocouple type, the terminal for thermocouple without thermocouple can detect present value.(In the case of RTD, it can be checked by a resistor tester)
SYS.E |  |  | - The internal parameter error - When the input side is affected by noise - System error in EEProm |  | Please contact Hanyoung Nux.
ADE,DDE |  |  | 250Ω resistors or internal AC/ DC converter problems in 4-20mA input |  | 1. Reboot 2. Remove the resistor 3. Remove the sensor 4. Check the error code and follow the corresponding troubleshooting again

## Slide 14: 4. Digital Temperature Controller
HY-48D
HY-72D
HY-8000
HY-8200S
4. Digital Temperature Controller
Error Code |  |  | Error Description |  | Troubleshooting
OVR / -OVR |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring

## Slide 15: 5. Temperature Indicator
BK6
RT9
5. Temperature Indicator
Error Code |  |  | Error Description |  | Troubleshooting
OVR / -OVR |  |  | Disconnection from inputting Voltage or current PV exceeding the maximum range of temperature |  | 1. Check input sensor type and code 2. Check input scale min/max value 3. Check sensor wiring

## Slide 16: 6. Counter/Timer
GE/GF series
6. Counter/ Timer
Error Code |  | Error Description |  | Troubleshooting
- |  | Reset button error |  | 1. Reset button is pressed continuously 2. If the reset button does not return, please cal for repair